Aliveness Test Procedure

for the

Magnetometer Instrument (MAG)

on the

Advanced Composition Explorer (ACE)

Submitted by the

Bartol Research Institute

University of Delaware

Newark, DE 19716

Initial Release: April 23, 1996

Revision A: April 29, 1996

Approvals:

Prepared by:		Date:

Jacques L’Heureux

ACE/MAG Instrument Manager

Bartol Research Institute

Approved by		Date:

Mario H. Acuña

ACE/MAG Instrument Scientist

Goddard Space Flight Center

Approved by		Date:

Norman F. Ness

ACE/MAG Experiment Manager

Bartol Research Institute

Document Change Log

| Revision | Date | Change Description | Prepared by |
| --- | --- | --- | --- |
| Initial Release | April 23, 1996 | Initial Release | Jacques L’Heureux |
| Rev. A | April 29, 1996 | Correction to Acceptable range for sensor readings | Jacques L’Heureux |

Table of Contents

PURPOSE

SIDE A or B OPERATIONS

1. SETUP

2. RECORD & VERIFY

3. CAL COMMAND ON

4. CAL COMMAND OFF

5. DEFAULT CONFIGURATION COMMAND

6. SHUTDOWN

# PURPOSE

This procedure will be used:

- during integration (with modifications) using the ACE GSE. Command will be sent by the S/C integrator from the ACE GSE and results will be gathered from the ACE GSE and in some cases, from the MAG GSE looking at telemetry information from the ACE GSE. The final procedure for this will be a modification of what is in this document.

# SIDE A or B OPERATIONS

## 1.	SETUP

1.	Enter DATE...........................................____________

2.	Enter Instrument Serial #..............................._________

3.	Verify sensor cables are connected......................_________

4.	Turn MAG power on to side A or B........................_________

5.	Turn 'Record Data' On; enter name of file........._______________

6.	Are the sensors in their shields? ................._______________

## 2.	RECORD & VERIFY

1.	RECORD Instrument Current..........................______________

(Acceptable Range 10 to 90 mA)

2.	RECORD Power Converter Temperature.................______________

(Acceptable Range -5 to +45 C)

3.	RECORD MAG A Range value and 'Auto/Manual' Mode....  ..._________

(Acceptable Range R0 to R7, Auto or Manual)

4.	RECORD MAG B Range value and 'Auto/Manual' Mode...  ...._________

(Acceptable Range R0 to R7, Auto or Manual)

5.	RECORD Instrument operating mode (0,1, or 2)............_________

(Acceptable Range 0)

6.	Verify FFT Graph O/K...................................._________

7.	RECORD MAG readings:

A SENSOR - X...__________ B SENSOR - X...__________

A SENSOR - Y...__________ B SENSOR - Y...__________

A SENSOR - Z...__________ B SENSOR - Z...__________

(Acceptable Range 001 to FFE Hex)

If any of the above readings are outside the acceptable range, execute SHUTDOWN (section 6), wait 30 seconds and restart at 1.4. Do not do this more than once.

## 3.	CAL COMMAND ON

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR - X...__________ B SENSOR - X...__________

A SENSOR - Y...__________ B SENSOR - Y...__________

A SENSOR - Z...__________ B SENSOR - Z...__________

(Acceptable Range 001 to FFE Hex)

## 4.	CAL COMMAND OFF

1.	SEND A MINOR MODE CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

- MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR - X...__________ B SENSOR - X...__________

A SENSOR - Y...__________ B SENSOR - Y...__________

A SENSOR - Z...__________ B SENSOR - Z...__________

(Acceptable Range 001 to FFE Hex)

## 5.	DEFAULT CONFIGURATION COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

Readout Mode to M0, All Else to Default State

Do this by sending the command: M0_CNFG_DF

2.	Verify status :- A & B CAL and EF are OFF, Auto Ranging is in effect, Mode = 0

3.	Acquire Data as needed.

## 6.	SHUTDOWN

1.	Turn OFF Data Record......................................________

2.	Turn OFF MAG Power........................................________