Bench Acceptance Test Procedure

for the

Magnetometer Instrument (MAG)

on the

Advanced Composition Explorer (ACE)

Submitted by the

Bartol Research Institute

University of Delaware

Newark, DE 19716

Initial Release: April 23, 1996

Revision A: April 29, 1996.

Approvals:

Prepared by:		Date:

Jacques L’Heureux

ACE/MAG Instrument Manager

Bartol Research Institute

Approved by		Date:

Mario H. Acuña

ACE/MAG Instrument Scientist

Goddard Space Flight Center

Approved by		Date:

Norman F. Ness

ACE/MAG Experiment Manager

Bartol Research Institute

Document Change Log

| Revision | Date | Change Description | Prepared by |
| --- | --- | --- | --- |
| Initial Release | April 23, 1996 | Initial Release | Jacques L’Heureux |
| Rev. A | April 29, 1996 | Correction to acceptable range for sensor readings | Jacques L’Heureux |
|  |  | Update of Data Items Mnemonics |  |

Table of Contents

PURPOSE

SIDE A OPERATIONS

A1. SETUP

A2. RECORD & VERIFY

A3. CAL COMMAND ON

A4. CAL COMMAND OFF

A5. CAL ON / MODE 1 COMMAND

A6. CAL OFF / MODE 1 COMMAND

A7. CAL ON / MODE 2 COMMAND

A8. CAL OFF / MODE 2 COMMAND

A9. EF ON COMMAND

A10. EF OFF COMMAND

A11. HEATER POWER

A12. SHUTDOWN

SIDE B OPERATIONS

B1. SETUP

B2. RECORD & VERIFY

B3. CAL COMMAND ON

B4. CAL COMMAND OFF

B5. CAL ON / MODE 1 COMMAND

B6. CAL OFF / MODE 1 COMMAND

B7. CAL ON / MODE 2 COMMAND

B8. CAL OFF / MODE 2 COMMAND

B9. EF ON COMMAND

B10. EF OFF COMMAND

B11. HEATER POWER

B12. SHUTDOWN

DATA ITEMS MNEMONICS, LOCATIONS AND SCALE FACTORS

# PURPOSE

This procedure will be used:

- before shipping at GSFC using the ACE simulator and the MAG GSE. Commands will be sent from the MAG GSE and results will be obtained from the MAG GSE

- after shipping at APL using the ACE simulator and the MAG GSE Commands will be sent from the MAG GSE and results will be obtained from the MAG GSE

during integration (with modifications) using the ACE GSE. Command will be sent by the S/C integrator from the ACE GSE and results will be gathered from the ACE GSE and in some cases, from the MAG GSE looking at telemetry information from the ACE GSE. The procedure for this will be a modification of what is in this document.

# SIDE A OPERATIONS

## A1.	SETUP

1.	Enter DATE...........................................____________

2.	Enter Instrument Serial #..............................._________

3.	Verify sensor cables are connected......................_________

4.	Turn MAG power on to side A............................._________

5.	Turn 'Record Data' On; enter name of file........._______________

6.	Are the sensors in their shields? ................._______________

## A2.	RECORD & VERIFY

1.	RECORD Instrument Current..........................______________

(Acceptable Range 10 to 90 mA)

2.	RECORD Power Converter Temperature.................______________

(Acceptable Range -5 to +45 C)

3.	RECORD MAG A Range value and 'Auto/Manual' Mode........_________

(Acceptable Range R0 to R7, Auto or Manual)

4.	RECORD MAG B Range value and 'Auto/Manual' Mode........_________

(Acceptable Range R0 to R7, Auto or Manual)

5.	RECORD Instrument operating mode (0,1, or 2)............_________

(Acceptable Range 0)

6.	Verify FFT Graph O/K...................................._________

7.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

If any of the above readings are outside the acceptable range, execute SHUTDOWN (section A12.2), wait 30 seconds and restart at A1.4. Do not do this more than once.

## A3.	CAL COMMAND ON

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A4.	CAL COMMAND OFF

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

- MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A5.	CAL ON / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A6.	CAL OFF / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

4.	Verify FFT Graph O/K................. .................._________

## A7. CAL ON / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A8.	CAL OFF / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A9.	EF ON COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF ON, MANUAL RANGE 7

B.  B EF ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_FLIP_R7

2.	Verify status :- A & B EF is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A10.	EF OFF COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF OFF, MANUAL RANGE 7

B.  B EF OFF, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B EF is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A11.	HEATER POWER

1.	Turn ON Heater Power to B & A.........................________

2.	RECORD the following:

A Power...___________       B Power...___________

(Acceptable value Less than 1.10 W)

A Temp....___________       B Temp....___________

(Acceptable Range -25 to +55 C)

3.	Turn OFF Heater Power to B & A...................________

4.	Verify Heater is OFF.............................________

## A12.	SHUTDOWN

1.	Turn OFF Data Record......................................________

2.	Turn OFF MAG Power........................................________

3.  Switch both power and telemetry interface cables at the

GSE Simulator box to side B.............................._________

# SIDE B OPERATIONS

## B1.	SETUP

1.	Verify sensor cable is connected........................_________

2.	Turn MAG power on to side B............................._________

3.	Turn 'Record Data' On; enter name of file........._______________

## B2.	RECORD & VERIFY

1.	RECORD Instrument Current..........................______________

(Acceptable Range 10 to 90 mA)

2.	RECORD Power Converter Temperature.................______________

(Acceptable Range -5 to +45 C)

3.	RECORD MAG B Range value and 'Auto/Manual' Mode....  ..._________

(Acceptable Range R0 to R7, Auto or Manual)

4.	RECORD MAG B Range value and 'Auto/Manual' Mode...  ...._________

(Acceptable Range R0 to R7, Auto or Manual)

5.	RECORD Instrument operating mode (0,1, or 2)............_________

(Acceptable Range 0)

6.	Verify FFT Graph O/K...................................._________

7.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

If any of the above readings are outside the acceptable range, execute SHUTDOWN (section B12.2), wait 30 seconds and restart at B1.2. Do not do this more than once.

## B3.	CAL COMMAND ON

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B4.	CAL COMMAND OFF

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

- MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B5.	CAL ON / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B6.	CAL OFF / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

4.	Verify FFT Graph O/K................. .................._________

## B7. CAL ON / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B8.	CAL OFF / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B9.	EF ON COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF ON, MANUAL RANGE 7

B.  B EF ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_FLIP_R7

2.	Verify status :- A & B EF is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B10.	EF OFF COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF OFF, MANUAL RANGE 7

B.  B EF OFF, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B EF is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B11.	HEATER POWER

1.	Turn ON Heater Power to B & A.........................________

2.	RECORD the following:

A Power...___________       B Power...___________

(Acceptable value Less than 0.75 W)

A Temp....___________       B Temp....___________

(Acceptable Range -25 to +55 C)

3.	Turn OFF Heater Power to B & A........................________

4.	Verify Heater is OFF.............................________

## B12.	SHUTDOWN

1.	Turn OFF Data Record......................................________

2.	Turn OFF MAG Power........................................________

# DATA ITEMS MNEMONICS, LOCATIONS AND SCALE FACTORS

| Data Item | Mnemonic | Telemetry Location | Scaling | Units |
| --- | --- | --- | --- | --- |
| A Power On/Off | MAG_PWR_A |  |  | ON/OFF |
| B Power On/Off | MAG_PWR_B |  |  | ON/OFF |
| Sensors Heater On/Off | MAG_INT_HTR |  |  | ON/OFF |
| Instrument Current | MAG_OR_SURVHTR_I |  | TBD | ma |
| A Sensor Heater Power | MAG_A_HTR_T |  | TBD | ma |
| B Sensor Heater Power | MAG_B_HTR_T |  | TBD | ma |
| A Sensor Temperature | MAG_A_T |  | TBD | DEG C |
| B Sensor Temperature | MAG_B_T |  | TBD | DEG C |
| Power Converter Temperature (8 bits) | PCTEMP | mf-6 bits 296 to 303 | TBD | DEG C |
| Instrument Current Monitor (8 bits) | CMON | mf-7 bits 296 to 303 | TBD | ma |
| MAG A Range Value (0-7) | ST2_R | mf-1 or mf-9, bits 299 to 301 |  | 0-7 |
| MAG B Range Value (0-7) | ST1_R | mf-0 or mf-8, bits 299 to 301 |  | 0-7 |
| MAG A Auto/Manual Mode (0/1) | ST2_A | mf-1 or mf-9, bit 298 |  | 0/1 |
| MAG B Auto/Manual Mode (0/1) | ST1_A | mf-0 or mf-8, bit 298 |  | 0/1 |
| Instrument Operating Mode(0,1,2) | ST2_M | mf-1 or mf-9, bits 296 and 297 |  | (HEX) |
| Primary Sensor X Reading (any mode) | PX | Any mf, bits 0 to 11 |  | (HEX) |
| Primary Sensor Y Reading (any mode) | PY | Any mf, bits 12 to 23 |  | (HEX) |
| Primary Sensor Z Reading (any mode) | PZ | Any mf, bits 24 to 35 |  | (HEX) |
| Secondary Sensor - Reading (Mode 0 or 1) | SX | Any mf, bits 180 to 191 |  | (HEX) |
| Secondary Sensor Y Reading (Mode 0 or 1) | SY | Any mf, bits 192 to 203 |  | (HEX) |
| Secondary Sensor - Reading (Mode 0 or 1) | SZ | Any mf, bits 204 to 215 |  | (HEX) |
| Secondary Sensor X Reading (Mode 2) |  | Not available |  |  |
| Secondary Sensor - Reading (Mode 2) |  | Not available |  |  |
| Secondary Sensor- Reading (Mode 2) |  | Not available |  |  |
| MAG A Cal On/Off (1/0) | ST2_C | mf-1 and mf-9, bit 302 | ON=1 | ON/OFF |
| MAG B Cal On/Off (1/0) | ST1_C | mf-0 and mf-8, bit 302 | ON=1 | ON/OFF |
| A Electronic Flipper On/Off | ST2_EF | mf-1 and mf-9, bit 303 | ON=1 | ON/OFF |
| B Electronic Flipper On/Off | ST1_EF | mf-0 and mf-8, bit 303 | ON=1 | ON/OFF |
| A Sensor is Primary if | ST1_S = 1 | mf-0 and mf-9, bit 296 |  |  |
| B Sensor is Primary if | ST1_S = 0 | mf-0 and mf-9, bit 296 |  |  |

mf = minor frame

First bit is bit 0

First minor frame is 0

Most significant bit is always first

Default state:	B = Primary, A =Secondary (ST1_S = 0)