Limited/Comprehensive

Performance Test Procedure

for the

Magnetometer Instrument (MAG)

on the

Advanced Composition Explorer (ACE)

Submitted by the

Bartol Research Institute

University of Delaware

Newark, DE 19716

DRAFT: April 29, 1996

Approvals:

Prepared by:		Date:

Jacques L’Heureux

ACE/MAG Instrument Manager

Bartol Research Institute

Approved by		Date:

Mario H. Acuña

ACE/MAG Instrument Scientist

Goddard Space Flight Center

Approved by		Date:

Norman F. Ness

ACE/MAG Experiment Manager

Bartol Research Institute

Document Change Log

| Revision | Date | Change Description | Prepared by |
| --- | --- | --- | --- |
| Draft | April 29,1996 | Draft | Jacques L’Heureux |

Table of Contents

SIDE A OPERATIONS

A1. SETUP

A2. RECORD & VERIFY

A3. CAL COMMAND ON

A4. CAL COMMAND OFF

A5. CAL ON / MODE 1 COMMAND

A6. CAL OFF / MODE 1 COMMAND

A7. CAL ON / MODE 2 COMMAND

A8. CAL OFF / MODE 2 COMMAND

A9. EF ON COMMAND

A10. EF OFF COMMAND

A11. HEATER POWER (COMPREHENSIVE TEST ONLY)

A12. SHUTDOWN

SIDE B OPERATIONS (COMPREHENSIVE TEST ONLY)

B1. SETUP

B2. RECORD & VERIFY

B3. CAL COMMAND ON

B4. CAL COMMAND OFF

B5. CAL ON / MODE 1 COMMAND

B6. CAL OFF / MODE 1 COMMAND

B7. CAL ON / MODE 2 COMMAND

B8. CAL OFF / MODE 2 COMMAND

B9. EF ON COMMAND

B10. EF OFF COMMAND

B11. HEATER POWER

B12. SHUTDOWN

DATA ACQUISITION (OPTIONAL)

DEFAULT CONFIGURATION COMMAND

SHUTDOWN

# SIDE A OPERATIONS

## A1.	SETUP

1.	Enter DATE...........................................____________

2.	Enter Instrument Serial #..............................._________

3.	Verify sensor cables are connected......................_________

4.	Turn MAG power on to side A............................._________

5.	Turn 'Record Data' On; enter name of file........._______________

6.	Are the sensors in their shields? ................._______________

## A2.	RECORD & VERIFY

1.	RECORD Instrument Current..........................______________

(Acceptable Range 10 to 90 mA)

2.	RECORD Power Converter Temperature.................______________

(Acceptable Range -5 to +45 C)

3.	RECORD MAG A Range value and 'Auto/Manual' Mode........_________

(Acceptable Range R0 to R7, Auto or Manual)

4.	RECORD MAG B Range value and 'Auto/Manual' Mode........_________

(Acceptable Range R0 to R7, Auto or Manual)

5.	RECORD Instrument operating mode (0,1, or 2)............_________

(Acceptable Range 0)

6.	Verify FFT Graph O/K...................................._________

7.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

If any of the above readings are outside the acceptable range, execute SHUTDOWN (section A12.2), wait 30 seconds and restart at A1.4. Do not do this more than once.

## A3.	CAL COMMAND ON

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A4.	CAL COMMAND OFF

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

- MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A5.	CAL ON / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A6.	CAL OFF / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

4.	Verify FFT Graph O/K................. .................._________

## A7. CAL ON / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A8.	CAL OFF / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A9.	EF ON COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF ON, MANUAL RANGE 7

B.  B EF ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_FLIP_R7

2.	Verify status :- A & B EF is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A10.	EF OFF COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF OFF, MANUAL RANGE 7

B.  B EF OFF, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B EF is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## A11.	HEATER POWER (COMPREHENSIVE TEST ONLY)

1.	Turn ON Heater Power to B & A.........................________

2.	RECORD the following:

A Power...___________       B Power...___________

(Acceptable value Less than 1.10 W)

A Temp....___________       B Temp....___________

(Acceptable Range -25 to +55 C)

3.	Turn OFF Heater Power to B & A...................________

4.	Verify Heater is OFF.............................________

## A12.	SHUTDOWN

1.	Turn OFF Data Record......................................________

2.	Turn OFF MAG Power........................................________

3.  Switch both power and telemetry interface cables at the

GSE Simulator box to side B.............................._________

# SIDE B OPERATIONS (COMPREHENSIVE TEST ONLY)

## B1.	SETUP

1.	Verify sensor cable is connected........................_________

2.	Turn MAG power on to side B............................._________

3.	Turn 'Record Data' On; enter name of file........._______________

## B2.	RECORD & VERIFY

1.	RECORD Instrument Current..........................______________

(Acceptable Range 10 to 90 mA)

2.	RECORD Power Converter Temperature.................______________

(Acceptable Range -5 to +45 C)

3.	RECORD MAG B Range value and 'Auto/Manual' Mode....  ..._________

(Acceptable Range R0 to R7, Auto or Manual)

4.	RECORD MAG B Range value and 'Auto/Manual' Mode...  ...._________

(Acceptable Range R0 to R7, Auto or Manual)

5.	RECORD Instrument operating mode (0,1, or 2)............_________

(Acceptable Range 0)

6.	Verify FFT Graph O/K...................................._________

7.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

If any of the above readings are outside the acceptable range, execute SHUTDOWN (section B12.2), wait 30 seconds and restart at B1.2. Do not do this more than once.

## B3.	CAL COMMAND ON

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B4.	CAL COMMAND OFF

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

- MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B5.	CAL ON / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B6.	CAL OFF / MODE 1 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 1

Do this by sending the command: M1_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 1.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

4.	Verify FFT Graph O/K................. .................._________

## B7. CAL ON / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL ON, MANUAL RANGE 7

B.  B CAL ON, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_CAL_R7

2.	Verify status :- A & B CAL is ON, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B8.	CAL OFF / MODE 2 COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A CAL OFF, MANUAL RANGE 7

B.  B CAL OFF, MANUAL RANGE 7

C.  MODE = 2

Do this by sending the command: M2_RANGE_7

2.	Verify status :- A & B CAL is OFF, Manual Range 7, Mode = 2.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B9.	EF ON COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF ON, MANUAL RANGE 7

B.  B EF ON, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_FLIP_R7

2.	Verify status :- A & B EF is ON, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B10.	EF OFF COMMAND

1.	SEND A CONFIGURATION COMMAND THAT INCLUDES THE FOLLOWING:

A.  A EF OFF, MANUAL RANGE 7

B.  B EF OFF, MANUAL RANGE 7

C.  MODE = 0

Do this by sending the command: M0_RANGE_7

2.	Verify status :- A & B EF is OFF, Manual Range 7, Mode = 0.

3.	RECORD MAG readings:

A SENSOR- X...__________ B SENSOR- X...__________

A SENSOR- Y...__________ B SENSOR- Y...__________

A SENSOR- Z...__________ B SENSOR- Z...__________

(Acceptable Range 001 to FFE Hex)

## B11.	HEATER POWER

1.	Turn ON Heater Power to B & A.........................________

2.	RECORD the following:

A Power...___________       B Power...___________

(Acceptable value Less than 0.75 W)

A Temp....___________       B Temp....___________

(Acceptable Range -25 to +55 C)

3.	Turn OFF Heater Power to B & A........................________

4.	Verify Heater is OFF.............................________

## B12.	SHUTDOWN

1.	Turn OFF Data Record......................................________

2.	Turn OFF MAG Power........................................________

# DATA ACQUISITION (OPTIONAL)

## DEFAULT CONFIGURATION COMMAND

1.	SEND the following command:  M0_CNFG_DF

2.	VERIFY status :- B & A CAL and EF are OFF, Automatic Ranging is in effect, Mode = 0.

3.	Acquire data as needed.

## SHUTDOWN

1.	Turn OFF Data Record......................................________

2.	Turn OFF MAG Power........................................________